## Sheet1
2017-04-13 | acer 2008 ¦ audio driver alc268 ¦ added wires connecting to line in ¦ uninstalled drivers
2017-04-14 | acer 2008 ¦ unable to recognise line in ¦ removed instergrated mic circuits ¦ solder bridges created between legs
2017-04-15 | Acer 2008 ¦ cleared solder bridges ¦ reinstalled linux ¦ removed capcitor and reference on mic input ¦ jack detect permenently on ¦ lost 5kΩ resistor
2017-04-16 | acer 2008 ¦ installed ubuntu ¦ merged c and D drive as max partions is 4 ¦ line in not working ¦ mic working with dc offset ¦ input circuitry in degisn ¦ initial resistr devidedor method would not work as load would change devider voltage ¦ opamp as voltage buffer possible solution ¦ all opamps not working ¦ laptop re assembled screws missing and too many of others
2017-06-06 | Soulution needed for reference voltage
2017-06-07 | Used inverting amplifier ¦ schotkey diode used for diode drop used for brign referece voltage to ground level
2017-06-08 | acer audio gain set to 0db ¦ volts per divition setting set at 17.4 ¦ seems to work well
2017-10-08 | simplifyed and redrafted circuitry for osciloscope ¦ using a inverting amplifyer with a voltage offset 'ground' of the refernce for the sound card ¦ works well with about 10mV accuracy ¦ idea to use voltage reference from the computer (ie running a wire from the inside) to get more stable votlage reference ¦ idea to have both a ac coupled and dc coupled amplifyer, an advanced feature would be for the dc offset be automoatically applied to the ac waveform, this would achieve greater resolution in the signal and also not loose the absolute data infomation
2017-12-12 | Created input protection though the use of fuses and clamping diodes ¦ created bom ¦ parts need total to about $25
2017-12-24 | Layed out pcb on PCB design ¦ devised set of rules to better design for breadboard ¦ added vref wire from laptop ¦ began to assemble circuit boards
2017-12-25 | Built input protection and built most of the input circuity ¦ realised problem were the capacitive inputs were all tied together to vref, solved by adding 1 MΩ bodge resistor between them. ¦ Checked continuity and fixed one error of a cold joint.
2017-12-26 | Added attachment test points to circuit ¦ computer drawn vref is necessary as it swings ± 0.2V ¦ issues with ground referencing and powering of opamps ¦ it is currently suspected that the low impedance 20 KΩ path of vref to ground and Vpp is causing it to collapse ¦ Α voltage follower or higher impedance path may be needed.
2017-12-27 | Read about op amp input latching which occurs when the input voltages excede the supply voltages ¦ issue with current circuit as the voyage reduction mechanism relies on inputs being greater than supplies ¦ may need to redo that part of the input which would constitute major modification as the circut is already assembled ¦ added common collector input vref amplifier to design to also be implemented in upcoming modifications ¦ with Lt spice test inconclusive it may not be an issue ¦ practical test is needed
2017-12-29 | Am able to (theoretical considered) have input voltages above the supply so long as the rate of change of voltage with time never exceeds that of the opamps devided by the gain. This is because of this is granted the opamps is able to maintain the voltage of the input pin and not allow it to swing above or below ¦ solved some differential equations to determine the correct capacitor size for the slew rate limiting ¦ this solution is going to be implemented ¦ added capacitors and vref common base amplifier ¦ tested circuit, issue with voltage reference being variable with supply voltage due to voltage devider ¦ another issue has with input s being very slow responding probably due to capacitor discharge ¦ may need to redesign the circuit to remove capacitor from circuit when not in use. ¦ Signal not seen in test though other problems may be preventing it ¦ vref is a very high impedance source, a 1MΩ resistor to ground will cause a significant voltage drop ¦ will need to use opamp voltage follower ¦ am going to redesign circuit to use one opamp per channel and switch capacitors in with switch ¦ and have oneopamp for a vref voltage buffer ¦ another solution would be to remove the 1 MΩ resistors between the capacitive input and ground ¦ the opamp should solve the issue of biasing by itself
2017-12-31 | Removed 1 MΩ resistors ¦ accedentally created short between vref and another pin ¦ no harm appears to have occurred¦ got working but the scale is not zeroed on zero volts ¦ this causes associated problems with scales and bias voltages ¦ soluton involves using voltage devider on the non inverting input of the inverting opamp ¦ the values of the resistors were derived mathematically ¦ this also requires a lower impedance vref and so an opampwill need to be used. ¦ The nessesity of all these design changes is going to require a redesign and the scalping of this one and the creation of a new board ¦ redesigned board to include all of the optimisation. ¦ one issue is that the circut had a permanent 39kΩ minimum impedance to AC signals to the requirements that adding Ac input coupling with a signal SPDT terminal switch
2018-01-01 | Desoldered components from old board
2018-01-02 | PCB layout for new schematic ¦ new design had fewer components than previous making the board smaller as well ¦ soldered circuit together¦ trick to get low wattage soldering iron to work is to wait untill it reheats after joint and apply flux when needed on difficult and bad surface joints
2018-01-03 | Drilled out isolation holes ¦ tested circuit for propped continuity ¦ no issues had ¦ soldered vref wore I. Wrong location and was wondering why circuit was not working ¦ desoldered to correct location ¦ the circut then worked on 0-5v with no need to change vref ¦ calibration amplitude about 30 V div ¦ adjusted v bias potentiometer to correct level ¦ good independence on supply voltage ¦ created list of components needed for propped installation ¦ created model of how all componets would fit in chasis
2018-01-04 | Modelled layout of components in box in blender ¦ possitoned all switches and front panel components
2018-04-08 | Marked out and drilled holes for components ¦ have decided to go for a solution which has the Vref and USB power perminantly attached cables and the audio connector as a plug on each end ¦ am going to attach a more permanent vref socket to the computer ¦ began wireing up the ratsness of internal cables and wires ¦ am using mostly enamle copper wire which is difficult to whett but otherwise good to work with ¦ soldering in small spaces it better with the small soldering iron though its power means that it takes time to heat up after each joint ¦ project comming on well
2018-04-09 | Finished wireing up oscilloscope ¦ added vref and power USB cable ¦ fit together nicely ¦ added vref port to laptop ¦ Needed to drill out and file around area of socket to be able to fit plug. ¦ tested oscilloscope on computer ¦ led channel works well, AC and DC ¦ right is not responsive to anything. ¦
2018-04-10 | Began trouble shooting right channel not working by testing continuity ¦ reflowing joint which looks to have bad connectivity ¦ found solution of output wire from opamp fell off resoldered back on ¦ issue solved apart from some order between the two channels.¦ Need to adjust trim pot to fix. ¦ Attempted to calibrate the trimpots ¦ for v bias grounded both inputs and adjusted untill the green right channel was at 0v on screen. ¦ Noted some temperature dependence ¦ need to swap left and right channel names on box as they were reversed. ¦ Adjusted amplification voltage by applying the same signal to both inputs and then adjusting the trimpots so that at all levels they had the same reading. ¦ Unable to do so at all voltage levels and so made the. Equal each other at 5v the error is a few percent. ¦ About 100mV off the actual Reading. ¦ Not very linear over a voltage Range ¦ not sure if it is sound card or circuits fault. ¦ screwed on back cover. ¦ AC coupling on right channel is not working ¦ need to secure all outside connectors better. ¦ Negative voltages work.
2018-04-11 | Feature that can be mounted upside down if the back plate is screwed in in the other way ¦ attached aluminium angle to backplate to allow mounting to desk ¦